## BGKD Report
E-BAM PLUS Background Calibration Report
Test Details |  |  | Test conditions & other notes
Serial Number: | W15096 |  | EBAM moved from room 419 to rooftop.
Location: | Paccar ETB roof lab |  | Sampling through mnfctr zero-assm but not
Performed by: | Patrick O'Keeffe |  | cyclone or PM10 inlets.
Email: | pokeeffe@wsu.edu
Start Time: | 2019-06-28 18:00 |  | This template contains example data to
End Time: | 2019-07-02 09:00 |  | demonstrate input formatting. Update
BKGD as-found: | 0.001 |  | "Test Details" and then paste data into the
BKGD during test: | 0 |  | other worksheet. Results are calculated below.
Results (mg/m^3) |  |  | Printing the entire workbook is recommended
72-hr mean: | -0.0017361 | No alarms | for final report storage (2 Letter-sized pages).
Hourly stdev (σ): | 0.0041666 | In spec
New BKGD value: | 0.0017361
New detect limit: | 0.0083333 |  | Form rev. 2019-07-10, CC-BY 4.0

## Paste Data Here
| Instructions for this sheet: |  |  |  |  | Do not drag-and-drop to prevent corrupting formulas
1. Paste 72 hours of E-BAM PLUS data into shaded table (B6)
2. Split data set to columns, if needed, using Data > Text to columns > Delimited (comma)
All values below still in ug/m^3
Hrs | Time | ConcRT(ug/m3) | ConcHR(ug/m3) | Flow(lpm) | WS(m/s) | WD(Deg) | AT(C) | RH(%) | BP(mmHg) | FT(C) | FRH(%) | BV(V) | PM | Status | ZeroAvg | Zero+1σ | Zero-1σ | Zero+2σ | Zero-2σ
1 | 2019-06-29 10:00 | 3 | 2 | 16.6 | 0 | 0 | 22.1 | 39 | 696 | 31 | 16 | 12.5 | 0 | 0 | -1.7361 | 2.4305 | -5.9028 | 6.5972 | -10.0694
2 | 2019-06-29 11:00 | 0 | 1 | 16.6 | 0 | 0 | 23.4 | 33 | 696 | 32.7 | 14 | 12.5 | 0 | 0 | -1.7361 | 2.4305 | -5.9028 | 6.5972 | -10.0694
3 | 2019-06-29 12:00 | -4 | -2 | 16.6 | 0 | 0 | 24.2 | 27 | 696 | 33.7 | 11 | 12.5 | 0 | 0 | -1.7361 | 2.4305 | -5.9028 | 6.5972 | -10.0694
4 | 2019-06-29 13:00 | -4 | -8 | 16.6 | 0 | 0 | 24.8 | 29 | 695 | 33.9 | 12 | 12.5 | 0 | 0 | -1.7361 | 2.4305 | -5.9028 | 6.5972 | -10.0694
5 | 2019-06-29 14:00 | -2 | 2 | 16.6 | 0 | 0 | 26.5 | 25 | 695 | 35.7 | 11 | 12.5 | 0 | 0 | -1.7361 | 2.4305 | -5.9028 | 6.5972 | -10.0694
6 | 2019-06-29 15:00 | 3 | -2 | 16.7 | 0 | 0 | 26.3 | 25 | 695 | 36.1 | 10 | 12.5 | 0 | 0 | -1.7361 | 2.4305 | -5.9028 | 6.5972 | -10.0694
7 | 2019-06-29 16:00 | -7 | -6 | 16.6 | 0 | 0 | 26.2 | 26 | 695 | 35.8 | 11 | 12.5 | 0 | 0 | -1.7361 | 2.4305 | -5.9028 | 6.5972 | -10.0694
8 | 2019-06-29 17:00 | -7 | -5 | 16.7 | 0 | 0 | 25.9 | 26 | 694 | 35.8 | 10 | 12.5 | 0 | 0 | -1.7361 | 2.4305 | -5.9028 | 6.5972 | -10.0694
9 | 2019-06-29 18:00 | -5 | -6 | 16.6 | 0 | 0 | 25.9 | 24 | 694 | 35.8 | 10 | 12.5 | 0 | 0 | -1.7361 | 2.4305 | -5.9028 | 6.5972 | -10.0694
10 | 2019-06-29 19:00 | -1 | 2 | 16.6 | 0 | 0 | 24 | 31 | 694 | 34.6 | 11 | 12.5 | 0 | 0 | -1.7361 | 2.4305 | -5.9028 | 6.5972 | -10.0694
11 | 2019-06-29 20:00 | 1 | 0 | 16.7 | 0 | 0 | 21.6 | 40 | 694 | 31.4 | 15 | 12.5 | 0 | 0 | -1.7361 | 2.4305 | -5.9028 | 6.5972 | -10.0694
12 | 2019-06-29 21:00 | 3 | 3 | 16.7 | 0 | 0 | 19.8 | 45 | 695 | 29.2 | 17 | 12.5 | 0 | 0 | -1.7361 | 2.4305 | -5.9028 | 6.5972 | -10.0694
13 | 2019-06-29 22:00 | -3 | -3 | 16.7 | 0 | 0 | 17.7 | 56 | 695 | 27.5 | 20 | 12.5 | 0 | 0 | -1.7361 | 2.4305 | -5.9028 | 6.5972 | -10.0694
14 | 2019-06-29 23:00 | 0 | -2 | 16.7 | 0 | 0 | 16.1 | 62 | 695 | 25.5 | 22 | 12.5 | 0 | 0 | -1.7361 | 2.4305 | -5.9028 | 6.5972 | -10.0694
15 | 2019-06-30 | 1 | 5 | 16.6 | 0 | 0 | 14.6 | 73 | 695 | 24.2 | 25 | 12.5 | 0 | 0 | -1.7361 | 2.4305 | -5.9028 | 6.5972 | -10.0694
16 | 2019-06-30 01:00 | 4 | 2 | 16.7 | 0 | 0 | 14.2 | 77 | 695 | 22.9 | 28 | 12.5 | 0 | 0 | -1.7361 | 2.4305 | -5.9028 | 6.5972 | -10.0694
17 | 2019-06-30 02:00 | -6 | -5 | 16.6 | 0 | 0 | 13.2 | 81 | 695 | 22.1 | 29 | 12.5 | 0 | 0 | -1.7361 | 2.4305 | -5.9028 | 6.5972 | -10.0694
18 | 2019-06-30 03:00 | -5 | -8 | 16.6 | 0 | 0 | 13.6 | 78 | 695 | 22 | 28 | 12.5 | 0 | 0 | -1.7361 | 2.4305 | -5.9028 | 6.5972 | -10.0694
19 | 2019-06-30 04:00 | -5 | -3 | 16.7 | 0 | 0 | 13.3 | 77 | 695 | 21.9 | 28 | 12.5 | 0 | 0 | -1.7361 | 2.4305 | -5.9028 | 6.5972 | -10.0694
20 | 2019-06-30 05:00 | -4 | -3 | 16.7 | 0 | 0 | 14.1 | 67 | 695 | 21.5 | 27 | 12.5 | 0 | 0 | -1.7361 | 2.4305 | -5.9028 | 6.5972 | -10.0694
21 | 2019-06-30 06:00 | -5 | -11 | 16.6 | 0 | 0 | 13.8 | 73 | 695 | 21.5 | 28 | 12.5 | 0 | 0 | -1.7361 | 2.4305 | -5.9028 | 6.5972 | -10.0694
22 | 2019-06-30 07:00 | -6 | -3 | 16.6 | 0 | 0 | 18.5 | 52 | 695 | 25.3 | 23 | 12.5 | 0 | 0 | -1.7361 | 2.4305 | -5.9028 | 6.5972 | -10.0694
23 | 2019-06-30 08:00 | 1 | 1 | 16.6 | 0 | 0 | 23.2 | 38 | 695 | 30.8 | 17 | 12.5 | 0 | 0 | -1.7361 | 2.4305 | -5.9028 | 6.5972 | -10.0694
24 | 2019-06-30 09:00 | -6 | 1 | 16.6 | 0 | 0 | 25.5 | 27 | 695 | 34.8 | 12 | 12.5 | 0 | 0 | -1.7361 | 2.4305 | -5.9028 | 6.5972 | -10.0694
25 | 2019-06-30 10:00 | 4 | -2 | 16.6 | 0 | 0 | 25.5 | 29 | 695 | 35.5 | 11 | 12.5 | 0 | 0 | -1.7361 | 2.4305 | -5.9028 | 6.5972 | -10.0694
26 | 2019-06-30 11:00 | -2 | -2 | 16.6 | 0 | 0 | 25.3 | 27 | 695 | 34.4 | 12 | 12.5 | 0 | 0 | -1.7361 | 2.4305 | -5.9028 | 6.5972 | -10.0694
27 | 2019-06-30 12:00 | 1 | 0 | 16.6 | 0 | 0 | 26.6 | 25 | 694 | 34.9 | 11 | 12.5 | 0 | 0 | -1.7361 | 2.4305 | -5.9028 | 6.5972 | -10.0694
28 | 2019-06-30 13:00 | 4 | 5 | 16.6 | 0 | 0 | 27.9 | 22 | 694 | 36.4 | 10 | 12.5 | 0 | 0 | -1.7361 | 2.4305 | -5.9028 | 6.5972 | -10.0694
29 | 2019-06-30 14:00 | 1 | 2 | 16.6 | 0 | 0 | 28.9 | 22 | 694 | 37.9 | 9 | 12.5 | 0 | 0 | -1.7361 | 2.4305 | -5.9028 | 6.5972 | -10.0694
30 | 2019-06-30 15:00 | 3 | 3 | 16.7 | 0 | 0 | 28.8 | 22 | 693 | 38 | 9 | 12.5 | 0 | 0 | -1.7361 | 2.4305 | -5.9028 | 6.5972 | -10.0694
31 | 2019-06-30 16:00 | 4 | 2 | 16.6 | 0 | 0 | 29 | 21 | 693 | 37.7 | 9 | 12.5 | 0 | 0 | -1.7361 | 2.4305 | -5.9028 | 6.5972 | -10.0694
32 | 2019-06-30 17:00 | -5 | 0 | 16.7 | 0 | 0 | 28.6 | 20 | 693 | 37.6 | 9 | 12.5 | 0 | 0 | -1.7361 | 2.4305 | -5.9028 | 6.5972 | -10.0694
33 | 2019-06-30 18:00 | 3 | 0 | 16.6 | 0 | 0 | 28.1 | 22 | 693 | 37.6 | 9 | 12.5 | 0 | 0 | -1.7361 | 2.4305 | -5.9028 | 6.5972 | -10.0694
34 | 2019-06-30 19:00 | 3 | 4 | 16.7 | 0 | 0 | 26.6 | 26 | 693 | 36.4 | 10 | 12.5 | 0 | 0 | -1.7361 | 2.4305 | -5.9028 | 6.5972 | -10.0694
35 | 2019-06-30 20:00 | 2 | -1 | 16.7 | 0 | 0 | 23.8 | 31 | 693 | 33.7 | 12 | 12.5 | 0 | 0 | -1.7361 | 2.4305 | -5.9028 | 6.5972 | -10.0694
36 | 2019-06-30 21:00 | -4 | -7 | 16.7 | 0 | 0 | 21.8 | 37 | 693 | 31.5 | 14 | 12.5 | 0 | 0 | -1.7361 | 2.4305 | -5.9028 | 6.5972 | -10.0694
37 | 2019-06-30 22:00 | -7 | -2 | 16.6 | 0 | 0 | 19.1 | 49 | 693 | 29.3 | 17 | 12.5 | 0 | 0 | -1.7361 | 2.4305 | -5.9028 | 6.5972 | -10.0694
38 | 2019-06-30 23:00 | -3 | -7 | 16.6 | 0 | 0 | 17.9 | 56 | 693 | 27.3 | 20 | 12.5 | 0 | 0 | -1.7361 | 2.4305 | -5.9028 | 6.5972 | -10.0694
39 | 2019-07-01 | -6 | -3 | 16.7 | 0 | 0 | 17.1 | 56 | 693 | 25.7 | 21 | 12.5 | 0 | 0 | -1.7361 | 2.4305 | -5.9028 | 6.5972 | -10.0694
40 | 2019-07-01 01:00 | -2 | -4 | 16.6 | 0 | 0 | 19.6 | 43 | 693 | 26.2 | 19 | 12.5 | 0 | 0 | -1.7361 | 2.4305 | -5.9028 | 6.5972 | -10.0694
41 | 2019-07-01 02:00 | 5 | 10 | 16.6 | 0 | 0 | 16.3 | 60 | 693 | 25.6 | 21 | 12.5 | 0 | 0 | -1.7361 | 2.4305 | -5.9028 | 6.5972 | -10.0694
42 | 2019-07-01 03:00 | -1 | -4 | 16.7 | 0 | 0 | 14.1 | 71 | 693 | 23.2 | 25 | 12.5 | 0 | 0 | -1.7361 | 2.4305 | -5.9028 | 6.5972 | -10.0694
43 | 2019-07-01 04:00 | -2 | -5 | 16.6 | 0 | 0 | 13.5 | 74 | 693 | 22 | 27 | 12.5 | 0 | 0 | -1.7361 | 2.4305 | -5.9028 | 6.5972 | -10.0694
44 | 2019-07-01 05:00 | -6 | -1 | 16.7 | 0 | 0 | 13.8 | 71 | 693 | 22 | 27 | 12.5 | 0 | 0 | -1.7361 | 2.4305 | -5.9028 | 6.5972 | -10.0694
45 | 2019-07-01 06:00 | -3 | -9 | 16.7 | 0 | 0 | 15.5 | 64 | 693 | 22.7 | 27 | 12.5 | 0 | 0 | -1.7361 | 2.4305 | -5.9028 | 6.5972 | -10.0694
46 | 2019-07-01 07:00 | -9 | -9 | 16.6 | 0 | 0 | 19.2 | 50 | 693 | 26.4 | 22 | 12.5 | 0 | 0 | -1.7361 | 2.4305 | -5.9028 | 6.5972 | -10.0694
47 | 2019-07-01 08:00 | -1 | 3 | 16.6 | 0 | 0 | 21.9 | 40 | 693 | 30.5 | 17 | 12.5 | 0 | 0 | -1.7361 | 2.4305 | -5.9028 | 6.5972 | -10.0694
48 | 2019-07-01 09:00 | -6 | -5 | 16.6 | 0 | 0 | 23.1 | 35 | 693 | 31.7 | 15 | 12.5 | 0 | 0 | -1.7361 | 2.4305 | -5.9028 | 6.5972 | -10.0694
49 | 2019-07-01 10:00 | -2 | -7 | 16.7 | 0 | 0 | 24.8 | 30 | 693 | 33.1 | 13 | 12.5 | 0 | 0 | -1.7361 | 2.4305 | -5.9028 | 6.5972 | -10.0694
50 | 2019-07-01 11:00 | -6 | -3 | 16.6 | 0 | 0 | 26.5 | 28 | 693 | 34.9 | 12 | 12.5 | 0 | 0 | -1.7361 | 2.4305 | -5.9028 | 6.5972 | -10.0694
51 | 2019-07-01 12:00 | 5 | 4 | 16.6 | 0 | 0 | 28.2 | 23 | 693 | 36.7 | 10 | 12.5 | 0 | 0 | -1.7361 | 2.4305 | -5.9028 | 6.5972 | -10.0694
52 | 2019-07-01 13:00 | 0 | 5 | 16.7 | 0 | 0 | 28.8 | 21 | 692 | 38.4 | 9 | 12.5 | 0 | 0 | -1.7361 | 2.4305 | -5.9028 | 6.5972 | -10.0694
53 | 2019-07-01 14:00 | 2 | -4 | 16.6 | 0 | 0 | 29.8 | 20 | 692 | 39.3 | 9 | 12.5 | 0 | 0 | -1.7361 | 2.4305 | -5.9028 | 6.5972 | -10.0694
54 | 2019-07-01 15:00 | 0 | -1 | 16.7 | 0 | 0 | 29.9 | 19 | 691 | 39.8 | 8 | 12.5 | 0 | 0 | -1.7361 | 2.4305 | -5.9028 | 6.5972 | -10.0694
55 | 2019-07-01 16:00 | -3 | 0 | 16.6 | 0 | 0 | 29.6 | 19 | 691 | 38.8 | 8 | 12.5 | 0 | 0 | -1.7361 | 2.4305 | -5.9028 | 6.5972 | -10.0694
56 | 2019-07-01 17:00 | 5 | 2 | 16.6 | 0 | 0 | 28.9 | 22 | 690 | 38.7 | 9 | 12.5 | 0 | 0 | -1.7361 | 2.4305 | -5.9028 | 6.5972 | -10.0694
57 | 2019-07-01 18:00 | -5 | -3 | 16.7 | 0 | 0 | 26.8 | 27 | 691 | 36.6 | 11 | 12.5 | 0 | 0 | -1.7361 | 2.4305 | -5.9028 | 6.5972 | -10.0694
58 | 2019-07-01 19:00 | 4 | 1 | 16.7 | 0 | 0 | 26 | 29 | 690 | 35.2 | 12 | 12.5 | 0 | 0 | -1.7361 | 2.4305 | -5.9028 | 6.5972 | -10.0694
59 | 2019-07-01 20:00 | -4 | -4 | 16.7 | 0 | 0 | 23.4 | 38 | 691 | 33.1 | 14 | 12.5 | 0 | 0 | -1.7361 | 2.4305 | -5.9028 | 6.5972 | -10.0694
60 | 2019-07-01 21:00 | -7 | -10 | 16.7 | 0 | 0 | 20.1 | 46 | 692 | 28.9 | 18 | 12.5 | 0 | 0 | -1.7361 | 2.4305 | -5.9028 | 6.5972 | -10.0694
61 | 2019-07-01 22:00 | -6 | -7 | 16.7 | 0 | 0 | 19.3 | 51 | 693 | 27.2 | 20 | 12.5 | 0 | 0 | -1.7361 | 2.4305 | -5.9028 | 6.5972 | -10.0694
62 | 2019-07-01 23:00 | -5 | -8 | 16.7 | 0 | 0 | 17.8 | 64 | 693 | 26.1 | 24 | 12.5 | 0 | 0 | -1.7361 | 2.4305 | -5.9028 | 6.5972 | -10.0694
63 | 2019-07-02 | -7 | 0 | 16.6 | 0 | 0 | 16.9 | 72 | 693 | 25.1 | 27 | 12.5 | 0 | 0 | -1.7361 | 2.4305 | -5.9028 | 6.5972 | -10.0694
64 | 2019-07-02 01:00 | 6 | 4 | 16.6 | 0 | 0 | 16.4 | 76 | 693 | 25.1 | 27 | 12.5 | 0 | 0 | -1.7361 | 2.4305 | -5.9028 | 6.5972 | -10.0694
65 | 2019-07-02 02:00 | -2 | -8 | 16.7 | 0 | 0 | 15.4 | 83 | 692 | 24.2 | 29 | 12.5 | 0 | 0 | -1.7361 | 2.4305 | -5.9028 | 6.5972 | -10.0694
66 | 2019-07-02 03:00 | -3 | -1 | 16.7 | 0 | 0 | 14.2 | 89 | 692 | 23.1 | 31 | 12.5 | 0 | 0 | -1.7361 | 2.4305 | -5.9028 | 6.5972 | -10.0694
67 | 2019-07-02 04:00 | -4 | -2 | 16.6 | 0 | 0 | 13.5 | 93 | 692 | 22.2 | 33 | 12.5 | 0 | 0 | -1.7361 | 2.4305 | -5.9028 | 6.5972 | -10.0694
68 | 2019-07-02 05:00 | 0 | -2 | 16.6 | 0 | 0 | 13.4 | 92 | 692 | 21.6 | 33 | 12.5 | 0 | 0 | -1.7361 | 2.4305 | -5.9028 | 6.5972 | -10.0694
69 | 2019-07-02 06:00 | -6 | -3 | 16.6 | 0 | 0 | 15.1 | 81 | 692 | 22.5 | 32 | 12.5 | 0 | 0 | -1.7361 | 2.4305 | -5.9028 | 6.5972 | -10.0694
70 | 2019-07-02 07:00 | 0 | 1 | 16.6 | 0 | 0 | 16.4 | 73 | 693 | 25 | 27 | 12.5 | 0 | 0 | -1.7361 | 2.4305 | -5.9028 | 6.5972 | -10.0694
71 | 2019-07-02 08:00 | -3 | -1 | 16.6 | 0 | 0 | 17.5 | 68 | 693 | 26.1 | 25 | 12.5 | 0 | 0 | -1.7361 | 2.4305 | -5.9028 | 6.5972 | -10.0694
72 | 2019-07-02 09:00 | 1 | 2 | 16.6 | 0 | 0 | 18.2 | 63 | 693 | 27 | 23 | 12.5 | 0 | 0 | -1.7361 | 2.4305 | -5.9028 | 6.5972 | -10.0694