## Sheet1
Serial Number: | [5 empty] | Audit Results
Location: | [6 empty] | Pass, as found
Full name: | [6 empty] | Pass, with changes
Email: | [6 empty] | Does not pass
Date (yyyy/mm/dd):
Tape Status | Y | N | Leak Tests | As-found | As-left | In-spec:
Spool covers OK? |  |  | Basic: |  |  | < 1.0 lpm (NA=break)
New tape installed? |  |  | Adv. lower: |  |  | < 1.0 lpm
Amount remaining: | % |  | Adv. upper: |  |  | < 1.0 lpm
|  |  | Tape test | EBAM | CTS (Qa)
Routine service | Y | N | with tape:
Clean PM10 particle trap |  |  | w/o tape:
Clean PM10 drip jar |  |  | Diff. (with - w/o): |  |  | < 0.2 lpm
Clean PM10 bug screen
Clean PM10 funnel |  |  | Span Membrane Test |  |  | Analog Output Test
Clean PM2.5 particle trap |  |  | Measured value (Mass) |  |  | Ch# |  | EBAM | DMM
Clean PM2.5 grit pot |  |  | Expected value (Span)
Clean sample nozzle |  |  | Difference (Error) |  | %
Clean tape support vane |  |  |  |  | In-spec: ≤ 5%
|  | As-found |  | As-left, if changed
Sensor Audit |  | EBAM | CTS (Qa) | EBAM | CTS (Qa) | In-spec:
Air Temperature (Ta) | [5 empty] | ±2 ºC
Barometric Pressure (BP) | [5 empty] | ±10 mmHg
Sheath Flow Rate (Qa) | [5 empty] | 3.0 ± 0.12 lpm
Sample Flow Rate (Qa) | [5 empty] | 16.7 ± 0.67 lpm
Calibration Standard Details
Notes